IceCube

PMT HV Base Board

Engineering Requirements Document (ERD)

Version 2.2h (Draft)

Revisions:

Version	1.0 	W. Stroewe

1.1	A. Karle

1.2	A. Karle		July 29, 2002

2.0	N. Kitamura		August 15, 2002

2.1a	N. Kitamura		September 13, 2002

2.1b	N. Kitamura		September 16, 2002

2.1c	N. Kitamura		September 25, 2002

2.2	N. Kitamura		October 10, 2002

2.2a	N. Kitamura		October 24, 2002

2.2b	N. Kitamura		October 25, 2002

2.2c	N. Kitamura		November 15, 2002

2.2d	N. Kitamura		November 18, 2002

2.2e	N. Kitamura		November 21, 2002

2.2f	N. Kitamura		December 12, 2002

2.2g	N. Kitamura		December 16, 2002

2.2h	N. Kitamura		January 13, 2003

DRAFT

Table of Contents

Table of Contents	2

List of Figures	3

List of Tables	3

1	GENERAL	4

1.1	Scope	4

1.2	Purpose	4

1.3	Precedence	4

1.4	Responsibilities	4

1.5	Records	4

1.6	Units	4

1.7	Glossary and Acronyms List	4

1.8	References	6

2	FUNCTIONAL OVERVIEW	6

3	PERFORMANCE REQUIREMENTS	9

3.1	The HV PMT Supply	9

3.1.1	General	9

3.1.2	Dynode Chain Voltage Distribution	9

3.1.3	Damping Resistor Requirements	10

3.1.4	HV Control	10

3.1.5	Anode Current Sourcing Capability	11

3.1.6	Stability	11

3.1.7	Noise	11

3.2	Electrical	12

3.2.1	Power	12

3.2.2	Ground	12

3.2.3	Anode Signal Connection Requirements	13

3.2.4	PMT Mounting Holes Requirements	16

3.2.5	Digital Functionality Requirements	18

3.3	Physical	21

3.3.1	Definition	21

3.3.2	Overall size and shape requirements	21

3.3.3	Component placement	22

3.3.4	Excluded area	22

3.3.5	Minimum trace spacing requirements	22

3.3.6	Manual soldering compatibility	22

3.4	Environmental	23

3.4.1	Temperature Range	23

3.4.2	Pressure Range	23

3.5	Miscellaneous	23

3.5.1	Conformal coating	23

3.5.2	Silkscreen	23

Appendix 1	Design Notes	27

List of Figures

Figure 2.1	Functional overview of the PMT HV Base board.	7

Figure 3.1	Split ground configuration requirement	13

Figure 3.2	Anode signal coupling transformer signal definition (Illustration purpose only.  See text for correct winding requirements).	14

Figure 3.3	Plated-thru PMT mounting hole locations viewed from the top-side of the PMT HV Base circuit board.  The numerical labels associated with the holes mark the corresponding PMT pin number whose signal assignments are defined in Table 3.3.	16

Figure 3.4	Solder pad specification	17

Figure 3.5	Ribbon connector signal assignment	20

Figure 3.7	PSL Drawing No. 5549B020.  PMT HV Base Board dimensional and component placement requirements.  The figure identifies suggested locations for the ribbon cable connector, the coaxial cable attachment, and the clean ground wire attachment.  The PCB material thickness is for reference only.  (5549020B_e.pdf)	25

Figure 3.8	PSL Drawing No. 5549C021.  PMT HV Base Board component envelope definition. (5549021C_e.pdf)	26

List of Tables

Table 2.1	Summary of electrical connections requirements	8

Table 3.1	Dynode chain voltage distribution (“Dyn” denotes the n-th dynode or Dynode n.  “Fn” denotes the n-th focusing electrode.)	10

Table 3.3	PMT Pin Assignment	18

Table 3.4	Power ON/OFF signal assignment	20

Table 3.6	Ribbon connector signal assignment	21

# GENERAL

## Scope

This IceCube Engineering Requirements Document (ERD) specifies the physical, functional and performance requirements for the PMT High Voltage Base circuit board.

# Purpose

This requirement documentation shall be applicable to the development, prototyping, testing, and verification of the PMT High-Voltage Base circuit board.

## Precedence

In the event of a conflict between the provisions of this document and any prior IceCube documents, the provisions of this document shall govern.  Conflicts with other documents are resolved by the Change Control Board.

## Responsibilities

1.4.1	Physics/Engineering is responsible for writing and updating these requirements to ensure they are correct, complete and current.

1.4.2	Quality Assurance is responsible for ensuring this document and changes to it are properly reviewed, approved and maintained.

## Records

Records of initial review, approval and changes (Engineering Change Notices, ECN’s) in design shall be maintained according to the established processes.

Units

Weights and measures in this document are expressed in the MKS International System of Units (SI).

## Glossary and Acronyms List

ADC		Analog-to-Digital Converter

AWG		American Wire Gauge

cm			Centimeter

CMOS	Complementary Metal Oxide Semiconductor

CS0		Chip-select bit 1

CS1		Chip-select bit 0

DAC		Digital-to-Analog Converter

DAQ		Data Acquisition System

DC			Direct Current

DOM		Digital Optical Module

DOMMB	Digital Optical Module Main Board

ERD		Engineering Requirements Document

HV			High Voltage

Hz			Hertz

ID		Inside Diameter

IDC		Insulation Displacement Connector

IPC		Institute for Interconnecting and Packaging Electronic Circuits

k			Kilo (103)

kg			Kilogram

LED		Light-Emitting Diode

MKS		Meter-kilogram-second

M			Mega (106)

m			Meter

mA		Milliampere

MOSI		Master-Out-Slave-In

MISO		Master-In-Slave-Out

mV		Millivolt

mW		Milliwatt

n			Nano (10-9)

OD		Outside Diameter

OM		Optical Module

Pa			Pascal

PCB		Printed Circuit Board

PE			Photoelectron

pF		Pico Farad

PMT		Photomultiplier Tube

P/N		Part Number

PSL	Physical Sciences Laboratory, University of Wisconsin-Madison

P/V ratio	Peak-to-valley ratio

s, sec		Second

SCLK		Serial Clock

SI			Système International d’Unités

SMB		Sub-Miniature B

SPE		Single Photoelectron

TBD		To Be Determined

UL		Underwriters Laboratory

V			Volt

VDC		Volt DC

W			Watt

# References

- IceCube DOM Main Board – PMT HV Base Board Interface Requirements (Document No. 9000-0006)

- DOM Main Board Hardware Requirements (Document No. 9000-0007)

- PSL Drawing No. 5549B020 (PMT HV Base Circuit Board)

- PSL Drawing No. 5549C021 (PMT Envelope)

- IPC-2221 (Generic Standard on Printed Board Design)

# FUNCTIONAL OVERVIEW

The photo-multiplier tube high-voltage base (PMT HV Base) board is a modular PCB component to be integrated into each of the approximately 5000 optical modules (OM), containing a photo-multiplier tube (PMT), that will be deployed in the Antarctic deep-ice (below several kilometers) for scientific research purposes.  The PMT referred to in this document is a Hamamatsu Model 7081-02 with a nominal size of 10 inches (25.4 cm) in diameter and a nominal gain of 108.

The PMT HV Base board is required to function continuously without service over the entire twenty-year span of the research project under the deep-ice condition.  The operating temperature of the PMT HV Base board is a function of the deployment depth of the Optical Module in the ice, ranging roughly from –20C to –40C, whereas the operating pressure is that of the internal pressure of the OM, approximating the Antarctic ambient pressure of ~0.5 atm.  See 5.3 Environmental.

The PMT HV Base board has physical and electrical connections inside the OM with the photo-multiplier tube (PMT) and the digital optical module (DOM) main board (MB), the latter serving as the master controller of the entire OM.  Figure 2.1 depicts the functional relationship among the PMT HV Base board, the DOMMB and the PMT.  Table 2.1 summarizes the electrical connections between the PMT HV Base Board and the DOM Main Board.

The purpose of the PMT HV Base board is to facilitate the following functions:

- Generate a series of high-voltages for the individual dynodes, focusing electrodes and the anode of the PMT, using the power provided by the DOMMB.

- Transfer the anode signal pulses from the PMT to the DOMMB without distortion through a coaxial cable.

- Respond to the digital control commands issued by the DOMMB for power on/off and for the adjustment of the high voltages.

- Provide a digital reading of the high voltage to the DOMMB upon request.

- Provide digital board identification information to the DOMMB upon request.

The mechanical installation of the PMT HV Base board is accomplished by inserting the PMT lead pins into the plated-thru holes arranged on the PCB and soldering the pins to the annular pad associated with each plated-thru hole; this procedure also establishes the electrical connections between the PMT and the PMT HV Base board.

Detailed functional and performance requirements are specified in the rest of the document.

Figure 2.1	Functional overview of the PMT HV Base board.

Table 2.1	Summary of electrical connections requirements

| Connection method | Explanation | Section |
| --- | --- | --- |
| Plated-thru mounting holes | The board is physically mounted to the PMT by soldering the pins to these holes, which also makes electrical connections. | 5.2.4 |
| Coaxial RG-180B/U or equivalent | Connection between the secondary of the anode signal coupling transformer and the DOM main board. The board shall be delivered with one end of the coaxial cable attached to it. The other end of the coaxial cable requires an SMB type connector. | 5.2.3 |
| IDC Ribbon cable | Digital signals DC power Power & digital ground A female connector is required on board. | 5.2.5.2 |
| 0.52 mm2 (20 AWG) stranded wire | “Clean ground” connection. The board shall provide a wire pad. | 5.2.2.1.E |

# PERFORMANCE REQUIREMENTS

## The HV PMT Supply

### General

This subsection specifies electrical requirements applicable to the HV PMT supply portion of the PMT HV Base BoardElectrical.

#### Note on requirements alternatives

A set of alternative requirements, replacing the requirements defined in Section 5.1.3, shall be issued at a later date as a supplement to this ERD.   The vendor of the PMT HV Base board shall be appropriately notified by IceCube as to whether the present requirements or the said alternative requirements are to be enforced.

#### HV Generation

The method of HV generation shall be compatible with all the requirements stated in this document.  In particular, the electrical impedance of the voltage sources for the individual dynodes must be sufficiently low in order to meet the anode current sourcing capability (5.1.4).

#### Definition

- “First dynode voltage” shall refer to the voltage between the cathode and the first dynode of the PMT.

- “Anode voltage” shall refer to the voltage between the first dynode and the anode of the PMT.

Dynode Chain Voltage Distribution

#### Cathode potential

The PMT cathode shall be at ground potential.

#### Dynodes

The voltage across the successive dynode stages shall be according to the values specified in  in which the values are expressed in terms of a factor to be multiplied by the voltage across Dynode 1 (first dynode) and Dynode 2.

#### Focusing electrodes

The voltage for the focusing electrodes, denoted as F1 – F3, shall also be determined by the factor specified in  multiplied by the voltage across Dynode 1 and Dynode 2.

Note:  F1 and Dy1 are at the same potential.  F2 and F3 are at the same potential.

### Damping Resistor Requirements

A 100 (5% or better) resistor shall be present between each one of the last dynodes (Dy8, Dy9 and Dy10) and the corresponding high-voltage sources.

The said damping resistors shall be installed at locations easily accessible for the IceCube engineers to shunt or replace after the PMT HV Base board has been mounted on the PMT.

Table 3.2	Dynode chain voltage distribution (“Dyn” denotes the n-th dynode or Dynode n.  “Fn” denotes the n-th focusing electrode.)

| Interval | Voltage relative to Dy1 - Dy2 |
| --- | --- |
| Dy2 - Dy3 | 1.25 |
| Dy3 - Dy4 | 0.83 |
| Dy4 - Dy5 | 0.42 |
| Dy5 - Dy6 | 0.25 |
| Dy6 - Dy7 | 0.30 |
| Dy7 - Dy8 | 0.38 |
| Dy8 - Dy9 | 0.55 |
| Dy9 - Dy10 | 0.75 |
| Dy1 - F1 | 0.15 |
| Dy1 - F2 | 0 |
| Dy1 - F3 | 0.15 |

### HV Control

#### HV controllability requirements

##### First dynode voltage

- The first dynode voltage shall be set to the factory default value of 700 VDC.

- There shall be a provision for changing the said factory default value after delivery by the IceCube personnel to a value in the range of 600 to 800 VDC, using a readily-accessible and reliable method, such as installing or replacing a resistor.

##### Cathode-to-anode voltage

The cathode-to-anode voltage shall be adjustable at least over the range of 1000 to 2000 VDC by means of a suitable digital code written to the DAC residing on the PMT HV Base board.

#### HV monitoring requirement

There shall be a provision for monitoring the cathode-to-anode voltage as a digital code of an ADC reading transmitted to the DOM main board.

#### Digital interface

##### DAC resolution

The DAC used for setting the HV shall have a resolution of 12-bit.

##### ADC resolution

The ADC used for monitoring the HV shall have resolution of 12-bit.

##### Digital code vs HV

- The digital code for setting and monitoring the HV shall be in 12-bit unsigned straight binary with the digital value 000(hex) representing 0 V.

- The digital value and the corresponding HV value shall have a linear relationship at least in the voltage range specified in 5.1.3.1 with the slope of 0.5 V per bit.

Anode Current Sourcing Capability

The HV generator of the PMT HV Base board shall support the following current sourcing capability in the sense that the output voltage does not drop more than 10 V while producing the specified current:

- DC anode current of 12 nA at –40 C (deep-ice).

- DC anode current of 240 nA at room temperature (laboratory).

- Square-pulse anode current of 100 mA lasting for 1 sec.

Stability

The drift rate for the voltages supplied to the dynodes and the anode shall be less than 4 V / week during the regular in-ice operation.

### Noise

The ripple voltage observed at the output of the secondary of the anode signal-coupling transformer shall be no greater than 0.5mVpp when the output is terminated with a 100  resistor.

## Electrical

### Power

The PMT HV Base circuit board shall receive the electrical power from the DOM Main Board through the ribbon cable connector specified in 5.2.5.2.

The power source available to the PMT HV Base circuit board shall be a ±5 VDC voltage source with a voltage uncertainty of ±5%.

The maximum power dissipation of the PMT HV Base board shall be 300 mW.

### Ground

#### Split ground configuration

##### Definitions

The PMT HV Base board shall have two isolated ground planes as defined below (See Figure 3.2):

- Digital and power ground shall be referenced by the incoming power from the DOM Main Board; the digital control and monitor circuitry including the ADC and the DAC; and, the regulator and switching circuitry for HV generation.

- Clean analog ground shall be referenced by the voltage multiplier, dynode resistive divider, and regulator feedback circuitry.

##### Isolation resistance

The two ground planes defined above shall have a minimum isolation resistance of 10 M.

##### Stray capacitance

The stray capacitance between the two ground planes shall be less than 50 pF.

##### Jumper requirement

The PMT HV Base board shall have solder pads for a solderable jumper (a zero-ohm resistor) between the two ground planes.

  - The installation of the jumper is optional and shall be decided by the IceCube engineers after the delivery.

  - The PMT HV Base board shall delivered without the jumper installed.

##### Wire soldering pad requirement

The PMT HV Base Board shall have a wire soldering pad for the purpose of attaching a 20 AWG (0.52 mm2 conductor area) stranded wire for the “clean analog ground” connection to the DOM Main Board.

Figure 3.2	Split ground configuration requirement

).Anode Signal Connection Requirements

#### Signal coupling method

The PMT HV Base board shall employ a coaxial toroidal transformer in order to deliver the PMT signal pulses to the DOM main board.

##### Transformer signal definition

The transformer shall consist of a coaxial cable wound around a toroidal magnetic core.

- The center conductor and the shielding conductor of the coaxial winding shall provide the primary winding and the secondary winding of the transformer, respectively.

- The primary conductor and the secondary conductor accessible at one end of the coaxial winding shall be designated as the “dotted side” of the terminals (See Figure 3.3).

##### Transformer specification

- The transformer shall consist of a coaxial cable wound nineteen (19) times around a toroidal magnetic core.

- The said coaxial cable shall be RG-178/U or RG-178B/U with a Teflon inner dielectric and a Teflon outer jacket.

- The said toroidal magnetic core shall be Magnetics Model ZH-42206-TC (www.mag-inc.com) or equivalent.

- Adequate spacing between the windings shall be provided in order to insure high-voltage isolation integrity.

- There shall be an adequate means to hold the windings in place (such as a plastic plug pressed into the toroidal center, or a “belly-band” around the transformer).

##### Primary side requirements

- The primary side of the coaxial transformer shall be terminated with a 100 resister across the primary terminals (“back termination”).

- The “dotted” side of the primary terminal shall connect to the PMT anode terminal.

- The “un-dotted” side of the primary terminal shall connect to the source of the anode high-voltage.

##### Secondary side requirements

- The secondary side of the coaxial transformer shall be connected to the DOM Main Board using a coaxial medium.

- The said coaxial medium shall be RG-180B/U or a similar coaxial cable with a characteristic impedance of 100.

- The center conductor and the shield conductor of the said coaxial cable shall connect to the “dotted” side and the “un-dotted” side of the secondary terminal of the transformer, respectively.

Figure 3.3	Anode signal coupling transformer signal definition (Illustration purpose only.  See text for correct winding requirements).

##### Coaxial cable installation requirement

The PMT HV Base Board shall be delivered with the coaxial cable specified in 5.2.3.1.D installed.

- Electrical connections of the coaxial cable shall be accomplished by direct soldering.

- The said electrical connections shall not degrade when the cable is pulled with a maximum of 5kg of force in any directions.

- The length of the coaxial cable shall be 20 1cm.

- The end of the said coaxial cable not attached to the PCB shall have a right-angle, crimp-type SMB connector.

- The said SMB connector shall be AEP 2715-1521-004, gold plated coax crimp type; or, Sealectro 51-128-9511.

Overview

| PMT Mounting Holes Requirements Electrical connections requirements are summarized in |  |
| --- | --- |

#### Plated-thru PMT mounting holes and soldering pads

The PMT HV Base board shall have plated-thru holes specified in Figure 3.4 in order to make electrical connections to the PMT.

Each of the said plated-thru holes shall have an annular soldering pad in the manner specified in Figure 3.5.

The locations of the said plated-thru holes on the PCB are as specified in Figure 3.7.

#### Signal assignment to the PMT mounting holes

The electrical signal assignment to the plated-thru PMT mounting holes are as shown in Figure 3.4 and Table 3.3.

The PMT HV

Base board shall not be required to have holes to accommodate the pins that are assigned “No connection”.

Figure 3.4	Plated-thru PMT mounting hole locations viewed from the top-side of the PMT HV Base circuit board.  The numerical labels associated with the holes mark the corresponding PMT pin number whose signal assignments are defined in Table 3.3.

Figure 3.5	Solder pad specification

Table 3.3	PMT Pin Assignment

| Pin # | Signal name | Description |
| --- | --- | --- |
| 01 | NC | No connection |
| 02 | Dy1 | Dynode #1 |
| 03 | F3 | Focus #3 |
| 04 | NC | No connection |
| 05 | Dy3 | Dynode #3 |
| 06 | NC | No connection |
| 07 | Dy5 | Dynode #5 |
| 08 | Dy7 | Dynode #7 |
| 09 | Dy9 | Dynode #9 |
| 10 | P | Anode |
| 11 | NC | No connection |
| 12 | NC | No connection |
| 13 | NC | No connection |
| 14 | Dy10 | Dynode #10 |
| 15 | Dy8 | Dynode #8 |
| 16 | Dy6 | Dynode #6 |
| 17 | Dy4 | Dynode #4 |
| 18 | NC | No connection |
| 19 | Dy2 | Dynode #2 |
| 20 | F1 | Focus #1 |
| 21 | F2 | Focus #2 |
| 22 | NC | No connection |
| 23 | NC | No connection |
| 24 | K | Cathode |

### Digital Functionality Requirements

#### Digital signal standard

##### The digital signals (logic levels and voltages) between the PMT HV Base board and the DOM main board shall comply with the 3V CMOS signal standard.

##### The PMT HV Base Board shall not rely on the 5V-tolerance of the DOM Main Board when transmitting a signal.

#### The ribbon cable

The PMT HV Base Board shall have electrical connections with the DOM Main Board through a ribbon cable, defined in this section, for the following purposes:

- Power and ground connections

- Digital signal connections

##### The ribbon cable type

The ribbon cable shall be a 1.27mm-pitch flat IDC ribbon cable.

##### The ribbon connector type

The PMT HV Base board shall have a 2.54mm-pitch female IDC-type connector for the ribbon cable at the location specified in Figure 3.8.

The said connector shall be an AMP Micro-MaTch top-entry receptacle with locking (Tyco Electronics P/N 2-338068-0).

##### Signal duplication requirement

Each signal, ground and power in the ribbon cable shall have two connector pins allocated to it.

##### Ribbon connector signal assignment

The signal assignment to the ribbon connector pins shall be as defined in Table 3.5.

#### Chip-select (CS0, CS1)

The two chip-select signals, CS0 and CS1, shall be used in combination to select one of the following three digital devices residing on the PMT HV Base board:

- DAC: Digital-to-analog converter

- ADC: Analog-to-digital converter

- IDENT: Board identification device (5.2.5.6)

##### The assignment of the logic levels to CS0 and CS1 shall be determined by the vendor of the PMT HV Base board.

#### MOSI, MISO and SCLK signals

- DAC shall use MOSI and SCLK for data and serial clock, respectively.

- ADC shall use MISO and SCLK for data and serial clock, respectively.

- IDENT shall use one or more of MOSI, MISO and SCLK.

#### Power ON/OFF Control

The PMT HV Base board shall support a power ON/OFF control by the DOM Main Board through the ribbon cable.

##### ON/OFF control signal

The signal assignment for the power ON/OFF control shall be as shown in Table 3.4.

##### Power-up default

When the PMT HV Base board is turned on, the HV output shall be consistent with the DAC digital code of 0x000 (See 5.1.3.3.A).

Table 3.4	Power ON/OFF signal assignment

| Logic Level | Meaning |
| --- | --- |
| 0 | OFF |
| 1 | ON |

#### Digital Board Identification

The PMT HV Base board shall provide a unique digital board identification number (board ID) upon request from the DOM main board.

The said digital board ID device shall comply with the Dallas 1-Wire protocol to communicate with DOMMB serving as the bus master.

Figure 3.6	Ribbon connector signal assignment

Table 3.5	Ribbon connector signal assignment

| Pin # | Signal Name | Description |
| --- | --- | --- |
| 01 | DGND | Digital and power ground |
| 02 | SCLK | Serial clock |
| 03 | SCLK |  |
| 04 | MOSI | Master-out-slave-in |
| 05 | MOSI |  |
| 06 | MISO | Master-in-slave-out |
| 07 | MISO |  |
| 08 | DGND |  |
| 09 | CS0 | Chip-select bit 0 (See ) |
| 10 | CS0 |  |
| 11 | CS1 | Chip-select bit1 (See ) |
| 12 | CS1 |  |
| 13 | ON/OFF | Board enable/disable |
| 14 | ON/OFF |  |
| 15 | +5V | Main power (+) |
| 16 | +5V |  |
| 17 | DGND |  |
| 18 | DGND |  |
| 19 | -5V | Main power (-) |
| 20 | -5V |  |

Physical

### Definition

The “bottom side” of the PMT HV Base board shall refer to the side of the PCB from which the PMT leads are inserted.  The “top side” of the PMT HV Base board shall refer to the side opposite to the bottom side.  The terms “top view” and “bottom view” shall refer to the views from the top side and the bottom side of the PMT HV Base board, respectively.

Note:  The “ice top view” is a view of the Optical Module components in ice seen from the ice top.  For the purpose of the PMT HV Base board, the “ice top view” and the “top view” are synonymous.

### Overall size and shape requirements

The overall shape of the printed circuit board of the PMT HV Base shall be circular and no larger than 100 mm in diameter.  A greater diameter or a deviation from the circular outline for accommodation of special components such as connectors and cable harnesses shall be permitted provided that there is sufficient clearance between such components and the interior surface of the pressure vessel of the optical module.  The dimensional requirements are summarized in Figure 3.7 and Figure 3.8.  The volume constraints shall apply to both of the mating pieces of the ribbon connectors.

### Component placement

The components may be placed on either the top side or the bottom side of the PCB within the constraints of the component envelope, except for the following items, whose locations are specified in Figure 3.8:

- Anode signal coupling transformer (Bottom side)

- Coaxial cable for the anode signal connection

- Ribbon cable connector (Bottom side.)

- “Clean ground” connection wire pad

The following items shall be installed at locations where IceCube engineers can easily access for modification after the PMT HV Base board has been mounted on the PMT:

- Components for adjusting the first dynode voltage (5.1.3.1.A)

- Damping resistors (5.1.2)

- Solder pads for the optional jumper (5.2.2.1.D)

Excluded area

No components shall be mounted in the areas so specified in Figure 3.8.

### Minimum trace spacing requirements

In compliance with the circuit board trace layout rules specified for “B-4 External Conductors with Permanent Polymer Coating” in IPC-2221, §6.3 Electrical Clearance, the following conditions shall be met for both DC voltages and AC peak voltages:

- For voltage difference greater than 100 V and less than 300 V, the minimum trace spacing shall be 0.4 mm.

- For voltage difference greater than 300 V and less than 500 V, the minimum trace spacing shall be 0.8 mm.

- For voltage difference greater than 500 V, the minimum trace spacing shall be 0.8 mm plus 0.00305 mm per every volt exceeding 500 V.

### Manual soldering compatibility

The PCB shall be compatible with the increased temperature during the manual soldering of the PMT pins at the solder pads specified in 5.2.4.

## Environmental

Temperature Range

#### Operation

The continuous operating temperature of the PMT HV Base board shall be in the range of –40°C to +27°C.

#### Storage

The storage temperature for the PMT HV Base board shall be in the range of –55°C to +45°C.

#### Component selection

To the extent practical, all the components used for the PMT HV Base board shall meet the lowest operating temperature of –55C, as specified by the component manufacturer, where “practical” means that this requirement applies to all resistors, capacitors and diodes; PCB material; conformal coating; and any other components that are readily available for the operating temperature of –55C or lower.

The vendor of the PMT HV Base board shall supply IceCube with a list of components used that do not meet the –55C or lower operating temperature.

Pressure Range

The PMT HV Base circuit board shall be capable of operating continuously inside the pressure vessel with the sustained internal pressure of 40,000 Pa to 100,000 Pa.

## Miscellaneous

### Conformal coating

Conformal coating is required on both sides of the PCB.

#### Masking requirement

Areas where soldering will be performed after delivery shall be properly masked from the conformal coating.

### Silkscreen

Silkscreen marking is required on the top side of the PCB.

#### Items to be identified

The said marking shall at least identify the following items:

- Component(s) for setting the first dynode voltage

- Plated-thru holes for PMT connection

- Solderable jumper

- Clean ground wire pad

(Figure 3.7	See PLS Drawing 5549020B_e.pdf)

Figure 3.7	PSL Drawing No. 5549B020.  PMT HV Base Board dimensional and component placement requirements.  The figure identifies suggested locations for the ribbon cable connector, the coaxial cable attachment, and the clean ground wire attachment.  The PCB material thickness is for reference only.  (5549020B_e.pdf)

(Figure 3.8	See PLS Drawing 5549021C_e.pdf)

Figure 3.8	PSL Drawing No. 5549C021.  PMT HV Base Board component envelope definition. (5549021C_e.pdf)

Appendix 1	Design Notes

5.1.1.2	HV Generation

? Justification:  Reference to a Cockroft-Walton type generator in the earlier versions of this document has been removed in order to avoid unnecessary constraints on sound technical solutions.

Dynode Chain Voltage Distribution

Note:  The voltage ratios are in accordance with the specification for Hamamatsu PMT Model R7081-02.

5.1.3.1	HV controllability requirements

Justification:  The first dynode voltage mainly controls the peak-to-valley ratio (P/V ratio), whereas the anode voltage mainly controls the gain of the PMT.  Once the first dynode voltage is set to a value corresponding to an adequate P/V ratio (greater than 2.2), the voltage should not require readjustment.  It is possible to set the first dynode voltage to be sufficiently high so that the P/V ratio always exceeds the minimum required value of 2.2, regardless of the parameter variations among the PMTs.  The PMT (Hamamatsu R7081-02) has a nominal operating voltage of 460V for the first dynode voltage.  The manufacturer of the PMTs, however, confirms that the first dynode voltage may be as high as 800V without adverse effects.  See the e-mail exchange between Kael Hanson and Yuji Yoshizawa, Application and Sales Engineer of Hamamatsu Photonics Electron Tube Center.  The latest correspondence from Yoshizawa also confirms and guarantees that the cathode-to-anode voltage may be raised to 2000V, rather than previously-stated 1800V, thereby allowing a greater gain adjustment range (i.e., 2000V minus first-dynode voltage).

5.1.3.2	HV monitoring requirement

Justification:  It is useful to be able to read the exact state of the HV.  The ADC resolution should meet the DAC resolution.

5.1.3.3	Digital interface

Justification:  High-performance ADCs and DACs are readily available with 12-bit resolution.  A 12-bit resolution, corresponding to 0.5 volt per bit, is adequate for both monitoring and setting the HVs.

Error: Reference source not found

5.1.4	Anode Current Sourcing Capability

Justification:  The DC current requirement is obtained by assuming the PMT gain of 5E7, the average number of photoelectrons giving rise to the anode pulse of 1.5, and the noise rate of 1 kHz in deep-ice and 20 kHz at room temperature (worst case).  The pulse current requirement is meant to assure the dynamic range supporting all pulses that are realistically possible in the ice.  (The Hamamatsu PMT supports up to 70 mA of anode current and we don’t want the PMT Base board to be the bottleneck of any physical measurements.)

Stability

? Justification:  The gain-voltage relationship of the PMT is a power law of the form G ~ VB, where B ranges from 8 to 10, depending on the PMT.  Assuming the worst case with B=10, a 2% change in gain would require a voltage stability of dV/V = (1/B)(dG/G) = 0.2%, and consequently, a dV of several volt.

5.1.5	Noise

Justification: The rule of thumb being applied is that the trigger threshold for the system should be about 1/6th of the amplitude of an SPE (5mV), and that the systematic noise should be a small contribution to the triggering at that threshold.

5.2.2.1	Split ground configuration

Justification: The optimum ground connections for noise immunity need to be determined experimentally in conjunction with the DOM main board.  The split ground configuration is prerequisite for noise immunity.  The jumper requirement allows flexibility for experimentation.

5.2.3.1.C	Primary side requirements

Note:  It is vital that the anode signal connection is made with the correct signal polarity.

5.2.3.1.D	Secondary side requirements

Justification:  This RG-180B/U cable has a characteristic impedance of 95.  The output impedance of the transformer circuit is approximately 100, and the output pulse from the anode circuit suffers the least distortion when driving the same impedance.

5.2.3.1.E	Coaxial cable installation requirement

Justification: It is desirable to have the fewest number of solder connections after delivery.  The stripping and soldering of the thin coaxial cable is a potential quality issue.  Direct soldering is more preferable than using a connector for reliability and cost reasons.

5.2.4.2	Signal assignment to the PMT mounting holes

Note:  These connectors have a characteristic impedance of 75.

Justification:  The pin assignments are consistent with the pinouts for Hamamatsu PMT Models R7081-01 and R7081-02.  The plated-thru hole and solder pad dimensions are similar to those for the AMANDA 1999 Tube Base.  The PMT supplier (Hamamatsu) has agreed to deliver the PMT with the “No connection” pins cut short or removed.

5.2.5.2	The ribbon cable

Justification:  The 1mm-pitch cable is consistent with a 2mm-pitch ribbon connectors.  These connectors are more space-saving than the traditional 2.54mm-pitch counterpart.

5.2.5.6	Digital Board Identification

Note: A recommended device for this purpose is a Dallas Semiconductor DS2401.

5.2.7	Overall size and shape requirements

The terms ice top view and ice bottom view have been defined by PSL for the purpose of defining the OM assembly.

Justification:  Constraints on the size of the PMT HV Base board arise from the envelope of the Benthos glass sphere; the depth of the PMT placement in the optical gel; the envelope of the PMT itself; and, the position of the sphere penetrator.  The mounted height of the PMT HV Base board above the PMT has been chosen to be 20 mm.  The drawings provided are intended to make available a generous volume for the PMT HV Base board.  The 11-inch PMT has been ruled out due to size constraints in favor of the 10-inch PMT.  The 100mm-overall diameter is met by the current Iseg prototype.

5.2.10	Minimum trace spacing requirements

Justification:  These rules are necessary to ensure a reliable high-voltage performance.  The rules assume that the conductor traces are on the outer layers of the PCB and that a permanent polymer overcoat (post-assembly conformal coating) is present.

5.2.11	Manual soldering compatibility

Note:  Hamamatsu specifies the soldering conditions as less than 3 sec at 250C at 10mm from the glass.  The PCB may undergo a greater level of heating without exceeding the Hamamatsu spec, since the board is to be mounted 20mm from the PMT glass.

Temperature Range

Justification:  The ice-top temperature of –52ºC and deep-ice temperature of –40ºC are expected for operation.  The higher upper limit for the storage temperature takes care of possible situations during transportation.